Simple manual for Pt Sputter

28/Jun./2018 Toshihiro Hayakawa

Preparation

- Make sure that power switch at lower right is OFF.

- Open upper cover and take off glass chamber.

※Don’t contact seal area of glass chamber with anything.

Take care that glass chamber rolls over.

- Set sample on sample stage.

- Incline sample stage slightly.

(※In case of sample with a lot of asperity）

- Set glass chamber and close upper cover.

Film Formation by Pt Sputter

- Switch on power switch and emission starts automatically.

- Wait until pressure in chamber down under 20Pa.

- Set “coating time” and ”sputtering” current”

・Make sure that[Mode](a) is “Auto”.

・Push [DISPRA](b) and set “coating time” by pushing (c).

・Select “sputtering current” from (d).

- Decide ”coating time” and “sputtering current” from following graph arbitrarily .

- Turn on tab for rotation of sample stage.

- Conduct Pt sputtering.

Push [START](f) and start sputtering.

※Emission and sputtering start automatically as following step.

(For about 15 minutes)

- After sputtering, turn off tab for rotation of stage.

- Switch off power switch.

- Wait for becoming atmosphere pressure.

- Open upper cover and take off glass chamber. Take back tilt of sample stage and take out sample.

Finish

- Clean off glass chamber with ethanol and Kim Wipes.

- Make sure that there`re no dust.

- Set glass chamber and upper cover.