## Slide 1
①CPS must be set around 10,000 by changing the probe current
②DT must be set between 20% and 30% by changing the time constant
Time constant
Spectrum/image/map and line/multi points
Fig. 1 Start up
(In advance, you must switch off the power of the IR camera)

## Slide 2
①Select spectrum tab
Fig. 2 Spectrum
④Start acquisition of spectrum
⑤wait 10 sec or more
⑥Repress this button to stop
②Clear the previous spectrum
③Clear the previous identification
⑥ auto identification
⑦Show the quantity
Detail setting of identification
Controller of the size and area of the spectrum

## Slide 3
Raw spectrum data
Save spectrum as graphic file
Save spectrum data as csv format
How to save spectrum in spectrum mode

## Slide 4
Fig. 3 Image
④ Get SEM image
③Clear the previous identification
⑦ Auto identification
⑥ Show the quantity
Detail setting of identification
①Select image tab
②Clear the previous spectrum
④Start acquisition of spectrum
⑤Wait 10 sec or more
⑥Repress this button to stop
EXIT the SEM control by EDS
whole area/ limited area/point to be measured
Change the arrangement of the sreen

## Slide 5
Raw spectrum data
Save SEM image as graphic file
Save SEM image and
spectrum raw file
NOT csv file of spectrum
How to save spectrum in image mode

## Slide 6
Fig. 4 mapping and line
④ Get SEM image
②Clear the previous spectrum
③Clear the previous identification
⑧ Auto identification
Show the quantity
Detail setting of identification
①Select map/line tab
②Clear the previous spectrum
⑤Start acquisition of spectrum
⑥Wait 10 sec or more
⑦Repress this button to stop
EXIT the SEM control by EDS
whole area/ limited area/point to be measured
Arrange the displayed items
⑨Select mapping / line
⑩Start mapping / line

## Slide 7
After M start button is pressed
①File save dialog opens.
Enter file name header.
Press 保存(S) button.
② Comment input dialog opens.
Enter the comment in text box.
Press OK button.
Repress the M start button to stop
Mapping measurment
Mapping immediately stops
Mapping stops after current scan finished
Make data file (Remove the check when you stop scan
by unexpected reason)
After stopping the scan, the data is saved automatically.

## Slide 8
Fig. 5 detail setting of identification
Enter the name of element you want to add
Change the atomic number
delete the element from the element list
add the element from the candidate list
candidate list is shown when you indicate the any peaks on the spectrum
Identified elements list to be quantified or mapped
If checked, only α lines are displayed
Close the detail setting
Detail setting of identification

## Slide 9
How to change the mapping color
Press this to open the details
①Remove this check
④Set this check
②Select the element
to be changed color
③Select new color
This process must be done in advance of the start of mapping

## Slide 10
① Select “spectrum Tab”
② Open saved csv file to be conversed.
③ ファイル(F) → 名前をつけて保存(A)
Raw spectrum data
Save spectrum as graphic file
Save spectrum data as csv format
How to transfer raw spc file into csv file